## Slide 1: OTHER SENSORS

## Slide 2: OTHER SENSORS
There are many sensors other than level, pressure, flow, and temperature that may not be encountered on a day to day basis—such as position, force, smoke, and chemical sensors—but play an equally important part in process control in today’s high-technology industries and/or for operator protection.
These sensors will not be discussed in as much detail as the sensors already discussed.
Position, distance, velocity, and acceleration sensors
Rotation sensors using light and Hall effect sensors
Force, torque, load cells, and balances
Smoke detectors, gas, and chemical sensors in industry
Sound and light measurements
Sound and optical devices

## Slide 3: PosItIon and MotIon SensIng
Many industrial processes require both linear and angular position and motion measurements.These are required in robotics, rolling mills, machining operations, numerically controlled tool applications, and conveyers.
In some applications it is also necessary to measure speed, acceleration, and vibration.
Some transducers use position sensing devices to convert temperature and/or pressure into electrical units and controllers can use position sensing devices to monitor the position of an adjustable valve for feedback control.
Velocity or speed is the rate of change of position. This can be a linear measurement, i.e., feet per second (ft/s), meters per second (m/s), and so forth, or angular measurement, i.e., degrees per second, radians per second, rate per minute (r/m), and so forth.
Acceleration is the rate of change of speed, i.e., feet per second squared (ft/s2),meters per second squared (m/s2), and the like for linear motion, or degrees per second squared, radians per second squared, and the like, in the case of rotational motion.
Vibration is a measure of the periodic motion about a fixed reference point or the shaking that can occur in a process due to sudden pressure changes, shock, or unbalanced loading in rotational equipment. Vibration sensors are used to monitor the bearings in heavy rollers such as those used in rolling mills.

## Slide 4: PosItIon and MotIon SensIng
Absolute position is the distance measured with respect to a fixed reference point and can be measured whenever power is applied.
Incremental position is a measure of the change in position and is not referenced to a fixed point. This type of sensing can give very accurate positioning of one component with respect to another.
Rectilinear motion is measured by the distance traversed in a given time, velocity when moving at a constant speed, or acceleration when the speed is changing in a straight line.
Angular position is a measurement of the change in position of a point about a fixed axis measured in degrees or radians, where one complete rotation is 360° or 2p radians. The degrees of rotation of a shaft can be absolute or incremental. These types of sensors are also used in rotating equipment to measure rotation speed as well as shaft position and to measure torque displacement.
Angular motion is a measure of the rate of rotation. Angular velocity is a measure of the rate of rotation when rotating at a constant speed about a fixed point or angular acceleration when the rotational speed is changing.

## Slide 5: PosItIon and MotIon SensIng
Potentiometers are a convenient method of converting the displacement in a sensor to an electrical variable. The wiper or slider arm of a linear potentiometer can be mechanically connected to the moving section of a sensor. Where rotation is involved, a single or multiturn (up to 10 turns) rotational type of potentiometer can be used.
Linear variable differential transformers (LVDT) are devices that are used for measuring small distances and are an alternative to the potentiometer. The device consists of a primary coil with two secondary windings one on either side of the primary. A movable core when centrally placed in the primary will give equal coupling to each of the secondary coils. When an AC voltage is applied to the primary, equal voltages will be obtained from the secondary windings which are wired in series opposition to give zero output voltage. When the core is slightly displaced an output voltage proportional to the displacement will be obtained. These devices are not as cost effective as potentiometers but have the advantage of being noncontact. The outputs are electrically isolated, accurate, and have better longevity than potentiometers.
Figure. (a) LVDT with a movable core and three windings and (b) the secondary voltage versus core displacement for the connections shown.

## Slide 6: PosItIon and MotIon SensIng
Light interference lasers are used for very accurate incremental position measurements.
Monochromatic light (single frequency) can be generated with a laser and collimated into a narrow beam. The beam is reflected by a mirror attached to the moving object which generates interference fringes with the incident light as it moves. The wavelength of the light generated by a laser is about 5 × 10−7 m, so relative positioning to this accuracy over a distance of 1/2 to 1 m is achievable.
Ultrasonic, infrared, laser, and microwave devices can be used for distance measurement. The time for a pulse of energy to travel to an object and be reflected back to a receiver is measured, from which the distance can be calculated, i.e., the speed of ultrasonic waves is 340 m/s and the speed of light and microwaves is 3 × 108 m/s.
Ultrasonic waves can be used to measure distances from 1 to about 50 m, whereas the light and microwaves are used to measure longer distances.
If an object is in motion, the Doppler effect can be used to determine its speed. The Doppler effect is the change in frequency of the reflected waves caused by the motion of the object. The difference in frequency between the transmitted and reflected signal can be used to calculate the velocity of the object.

## Slide 7: PosItIon and MotIon SensIng
Hall effect sensors detect changes in magnetic field strength. The Hall effect occurs in semiconductor devices and is shown in Figure. Without a magnetic field the current flows directly through the semiconductor plate and the Hall voltage is zero. Under the influence of a magnet field, as shown, the current path in the semiconductor plate becomes curved, giving a Hall voltage between the sides adjacent to the input/output current.
Magneto resistive element (MRE) is an alterative to the Hall effect device. In the case of the MRE its resistance changes with magnetic field strength.

## Slide 8: PosItIon and MotIon SensIng
Optical devices detect motion by sensing the presence or absence of light. Figure shows two types of optical discs used in rotational sensing. Figure A shows an incremental optical shaft encoder. Light from the light-emitting diode (LED) shines through windows in the disc on to an array of photodiodes. As the shaft turns, the position of the image moves along the array of diodes. At the end of the array, the image of the next slot is at the start of the array. The relative position of the wheel with respect to its previous location can be obtained by counting the number of photodiodes traversed and multiplying them by the number of slots monitored.

## Slide 9: PosItIon and MotIon SensIng
Figure shows an absolute position encoder. An array of LEDs (one for each window) with a corresponding photo detector for each window can give the position of the wheel at any time. Only three windows are shown in the figure, for greater accuracy more slots would be used. The pattern shown on the disc is for the gray code.
Other patterns may be used on the disc such as the binary code. Optical devices have many uses in industry other than for the measurement of the position and speed of rotating equipment. Optical devices are used for counting objects on conveyer belts on a production line, measurement and control of the speed of a conveyer belt, location and position of objects on a conveyer, location of registration marks for alignment, bar code reading, measurement and thickness control, and detecting for breaks in filaments and so forth.
Power lasers can also be included with optical devices as they are used for scribing and machining of metals, laminates, and the like.

## Slide 10: PosItIon and MotIon SensIng
Accelerometers sense speed changes by measuring the force produced by the change in velocity of a known mass. These devices can be made with a cantilevered mass and a strain gauge for force measurement or can use capacitive measurement techniques. Accelerometers are now commercially available, made using micromachining techniques. The devices can be as small as 500 μm × 500 μm, so that the effective loading by the accelerometer on a measurement is very small. The device is a small cantilevered seismic mass that uses capacitive changes to monitor the position of the mass. Piezoelectric devices similar to the one shown in Figure are also used to measure acceleration. The seismic mass produces a force on the piezoelectric element during acceleration which causes a voltage to be developed across the element.
Accelerometers are used in industry for the measurement of changes in velocity of moving equipment, in the automotive industry as crash sensors for air bag deployment, and in shipping crates where battery operated recorders are used to measure shock during the shipment of expensive and fragile equipment. Vibration sensors typically use acceleration devices to measure vibration.

## Slide 11: Force,Torque, and Load Cells
Many applications in industry require the measurement of force or load.
Force is a vector and acts in a straight line, it can be through the center of a mass, or be offset from the center of the mass to produce a torque, or with two forces a couple.
Force can be measured with devices such as strain gages. In other applications where a load or weight is required to be measured the sensor can be a load cell.
Mass is a measure of the quantity of material in a given volume of an object.
Force is a term that relates the mass of an object to its acceleration and acts through its center of mass, such as the force required to accelerate a mass at a given rate.
Forces are defined by magnitude and direction and are given by the following:
Force (F ) = mass (m) × acceleration (a)

## Slide 12: Force,Torque, and Load Cells
Torque occurs when a force acting on a body tends to cause the body to rotate and is defined by the magnitude of the force times the perpendicular distance from the line of action of the force to the center of rotation.
Units of torque are pounds (lb), feet (ft), or newton meter (N⋅m).
Torque is sometimes referred to as the moment of the force, and is given by
Torque (t) = F × d

## Slide 13: Force and torque measurIng devIces
A couple occurs when two parallel forces of equal amplitude, but in opposite directions, are acting on an object to cause rotation, as shown in Figure and is given by the following equation:
Couple (c) = F × d
Analytical or lever balance is a device that is simple and accurate, and operates on the principle of torque comparison. Figure shows a diagram of a balance.

## Slide 14: Force and torque measurIng devIces
When in balance the torque on one side of the fulcrum is equal to the torque on the other side of the fulcrum, from which we get the following:
W1 × L = W2 × R
where W1 is a weight at a distance L from the fulcrum and W2 the counter balancing weight at a distance R from the fulcrum.
Spring transducer is a device that measures weight by measuring the deflection of a spring when a weight is applied, as shown in Figure.
The deflection of the spring is proportional to the weight applied (provided the spring is not stressed), according to the following equation:
F = Kd
where F = force in pounds or newtons
K = spring constant in pounds per inch or newtons per meter
d = spring deflection in inches or meters

## Slide 15: Force and torque measurIng devIces
Hydraulic and pneumatic devices can be used to measure force. This can be done by monitoring the pressure in a cylinder when the force (pounds or newtons) is applied to a piston as shown in Figure. The relation between force (F) and pressure ( p) is given by
F = pA
where A is the area of the piston.

## Slide 16: Force and torque measurIng devIces
Piezoelectric devices produce an electrical charge between the opposite faces of a crystal when the crystal is deformed by a force that makes them suitable for use as a force sensor. Many crystals exhibit the piezoelectric effect. Some common crystals are as follows:
Quartz
Rochelle salt
Lithium sulphate
Tourmaline

## Slide 17: Quartz devIces
Quartz devices have good sensitivity but have high output impedance. The output voltage drifts under low loading due to noise and temperature effects, but is well suited for measuring rapidly changing forces as well as static forces.
Tensile and compressive forces are measured with strain gauges; a strain gauge can use a piezoresistive material or other types of material that changes their resistance under strain. Figure shows the use of a strain gauge to measure the strain in a solid body under stress from a tensile force, in this case the material under tension elongates and narrows.
Strain gauges, as shown, are used to measure stress in a material from which the properties of the material can be calculated. A strain gauge can be used to measure stress from compressive forces as shown in Figure. An object under compressive forces will shorten and fatten and the strain can be measured.

## Slide 18: capacıtıve load cell
Weight measurements are made with load cells which can be capacitive, electromagnetic, use piezoelectric elements, or strain gauges. A capacitive load cell is shown in Figure.
The capacitance is measured between a fixed plate and a diaphragm. The diaphragm moves towards the fixed plate when force or pressure is applied, giving a capacitive change proportional to the force.

## Slide 19: Smoke and ChemIcal Sensors
The detection of smoke, radiation, and chemicals is of great importance in industrial processing not only as it relates to the safety of humans and the control of environment pollution both atmospheric and ground, but is also used in process control applications to detect the presence, absence, or levels of impurities in processing chemicals.
Smoke detectors and heat sensors (automatic sprinklers) are now commonplace in industry for the protection of people, equipment, and monitoring and control of chemical reactions.
Low-cost smoke detectors using infrared sensing or ionization chambers are commercially available. Many industrial processes use a variety of gases in processing–such as inert gases (nitrogen)—to prevent contamination from oxygen in the air, or conversely, gases or chemicals can be introduced to give a desired reaction.
It is therefore necessary to be able to monitor, measure, and control a wide variety of gases and chemicals. Awide variety of gas and chemical sensors are available. Of these, the Taguchi-type of sensor is one of the more common types of sensors.

## Slide 20: Smoke and chemIcal measurIng devIces
Infrared sensors detect changes in the signal received from a light emitting diode due to the presence of smoke in the light path or the presence of an object in the light path.
Ionization chambers are devices that detect the leakage current between two plates that have a voltage between them. The leakage occurs when carbon particles from smoke are present and provide a conductive path between the plates.
Taguchi-type sensors are used for the detection of hydrocarbon gases, such as carbon monoxide and dioxide, methane, and propane. The Taguchi sensor has an element coated with an oxide of tin that combines with hydrocarbon to give a change in electrical resistance which can be detected. Periodically the element is heated and the chemical reaction is reversed, reducing the coating back to tin oxide. Likewise, the sensing process can be repeated. The tin oxide can be made sensitive to different hydrocarbons by using different oxides of tin, different deposition techniques, and so on.

## Slide 21: Sound and LIght
The measurement of sound and light is important as it relates to the sense of hearing and sight, as well as many industrial applications.
Sound pressure waves can induce mechanical vibration and hence failure. Excessive sound levels produce noise pollution.
Light and its measurement is used in many industrial applications for high-accuracy linear measurements, location of overheating (infrared), object location and position measurements, photo processing, scanning, readers (bar codes), and so forth.
Sound waves are pressure waves that travel through air, gas, solids, and liquids, but cannot travel through space or a vacuum unlike radio (electromagnetic) waves.
X-rays should be mentioned at this point as they are used in the process control industry and are also electromagnetic waves.
X-rays are used primarily as inspection tools; the rays can be sensed by some light-sensing cells and can be very hazardous if proper precautions are not taken.

## Slide 22: Sound and lıght measurıng devıces
Microphones are transducers used to convert sound levels into electrical signals, i.e., electromagnetic, capacitance, ribbon, crystal, carbon, and piezoelectric microphones can be used. The electrical signals can then be analyzed in a spectrum analyzer for the various frequencies contained in the sounds or just to measure amplitude.
Sound level meter is the term given to any of the variety of meters for measuring and analyzing sounds.
Photocells are used for the detection and conversion of light intensity into electrical signals. Photocells can be classified as photovoltaic, photoconductive, photoemissive, and semiconductor.
Photovoltaic cells develop an emf in the presence of light. Copper oxide and selenium are examples of photovoltaic materials. A microammeter calibrated in lux (lm/m2) is connected across the cells and measures the current output.